# Automated Echocardiographic Measurement of Left Ventricular Systolic Function

Takeshi Hozumi, MD, Department of Internal Medicine and Cardiology,

Osaka City University School of Medicine, Osaka, Japan

The advanced automated contour tracking (AACT) method has been developed for accurate automated detection of the left ventricular (LV) endocardial boundary. LV ejection fraction (EF) can be estimated accurately by applying the AACT method to two orthogonal echocardiogrphic images in patients with ischemic heart disease (IHD) even when regional wall motion abnormalities exist. We evaluated 47 patients with known or suspected IHD. Biplane LVEF from apical 4- and 2-chamber views was measured off-line by the AACT method using disk summation method. In 41 of 47 patients (87.0%), automated tracing of the endocardial border was adequately achieved with the AACT method. LVEF measured by the AACT method was correlated well with that measured by quantitative gated SPECT (QGS) (y=0.97x+2.4, r =0.91). The mean difference between AACT and QGS was 0.6±5.5%. The mean time required for analysis of one set of images during one cardiac cycle by the AACT method was much shorter than that required by manual tracing method (7 ± 1 vs. 37 ± 4 sec, p<0.0001). Thus, the biplane AACT method provides accurate and quick measurement of LVEF in patients with IHD.

Applicattion of AACT method to apical 4- and 2-chamber views and long-axis view provides simplified automated three-dimensional (3D) measurement of LV volumes and EF. We examined  37 consecutive patients who underwent QGS because of suspected ischemic heart disease (IHD). In every patient, apical 4- and 2-chamber views and long-axis view were obtained by 2D echocardiography. In each case, three sample points was placed on both sides of the mitral annulus and the LV apex, in the end-diastolic image of three apical views. In the apical long-axis view, additional one points were placed on the basal septum in the end-diastolic (ED) image. In 30 of 37 patients (81%), adequate images were obtained for automated 3D LV volume and EF analyses. ED and end-systolic (ES) LV volumes and EF measured by the automated 3D method was correlated well with that by QGS (y=0.82x+10.4, r =0.97, y=0.84x+3.3, r =0.98 and y=0.89x+6.1, r =0.92, respectively). The mean differences in LV volumes and EF between automated 3D method and QGS were 6.4±12.4ml, 4.4±9.1ml and –0.1±6.0%, respectively. Thus, the simplified automated 3D method with AACT provides accurate measurement of LV volumes and EF in patients with suspected IHD.